Gamma ray spectroscopy with Ge(Li) and Abundance of Uranium Isotopes

ANALYSIS NOTE

The primary objective was to collect data of the energy spectrum for Cobalt-60, Cesium-137, Europium-152 and Barium-133. The tasks of the experiment are:

1) Finding the calibration formula and obtaining the relative detector efficiency.

2) Finding the Uranium abundance by calculating N(235)/N(238).

INSTRUMENTAL SPECIFICATIONS:

Lab gain settings:

- Fine gain: 1.195

- Course gain: 20

- Shaping: 3 microseconds

- Total = 23.9

DATA COLLECTED:

Table 1: Data collected for Cs, Co and Ba

| Element | Bin | Peak | Mean | Rms | Count | EOM | Precision | Gamma |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cesium-137 | 3558.0 | 263.0 | 3556.63 | 3.70 | 2302.0 | 0.08 | 2.08 | 661.657 |
| Cobalt -60 | 6202.0 | 145.0 | 6202.35 | 5.30 | 1508.0 | 0.14 | 2.58 | 1173.228 |
| Cobalt - 60 | 7026.0 | 120.0 | 7025.48 | 5.25 | 1246.0 | 0..15 | 2.83 | 1332.492 |
| Barium - 133 | 1563.0 | 115.0 | 1563.92 | 5.00 | 968.0 | 0.16 | 3.21 | 274.6 |
| Barium - 133 | 1701.0 | 257.0 | 1700.73 | 3.05 | 1792.0 | 0.07 | 2.36 | 302.851 |
| Barium - 133 | 1976.0 | 669.0 | 1976.66 | 2.90 | 4607.0 | 0.04 | 1.47 | 356.013 |

EQUATIONS USED:

Detector Efficiency=N_iNrefIref*(I_i)

Here N_i: The number of counts in the rest of the peaks, I_i: the corresponding relative intensities of the peaks.

Isotopic Ratio:

number of gammas emitted=number of gammas in the peakefficiency

number of decays of uranium =number of gammas emittedRelative Intensity

number of atoms of uranium=number of decays of uraniumλ

λ=log22t1/2

Isotopic ratio=number of atoms of U235number of atoms of U238

Percentage value=number of atoms of U235number of atoms of U235+number of atoms of U238

Energy Calibration:

DAY: 01

Europium-152

Figure 1: Europium peaks and corresponding gamma energy values

Using the bin count and the corresponding gamma values for Cesium, Cobalt and Barium, the Gamma energy vs Bin count graph was obtained. This graph was used to obtain the calibration fit function.

Figure 2: Energy calibration fit function

The calibration function is y = 0.193x – 26.015

The corresponding gamma ray energies for Eu-152 peaks:

Table 2: Gamma ray energies

| Literature | Experimental | Residual |
| --- | --- | --- |
| 121.000 | 120.350 | -0.6490 |
| 244.700 | 244.675 | -0.0248 |
| 344.300 | 344.443 | 0.1438 |
| 778.900 | 779.095 | 0.1951 |
| 964.079 | 964.131 | 0.0523 |
| 1085.869 | 1085.555 | -0.3137 |

DAY: 02

Europium-152

Figure 3: Europium peaks and corresponding gamma energy values

The calibration function is y = 0.193x – 26.015

The corresponding gamma ray energies for Eu-152 peaks:

Table 3: Gamma ray energies

| Literature | Experimental | Residual |
| --- | --- | --- |
| 121.000 | 120.350 | -0.6490 |
| 244.700 | 244.675 | -0.0248 |
| 344.300 | 344.443 | 0.1438 |
| 778.900 | 779.095 | 0.1951 |
| 964.079 | 964.518 | 0.4390 |
| 1085.869 | 1085.941 | 0.0729 |

Day 03:

Europium-152

Figure 4: : Europium peaks and corresponding gamma energy values

The calibration function is y = 0.193x – 26.015

The corresponding gamma ray energies for Eu-152 peaks:

Table 4: Gamma ray energies

| Literature | Experimental | Residual |
| --- | --- | --- |
| 121.000 | 120.157 | -0.8423 |
| 244.700 | 244.288 | -0.4115 |
| 344.300 | 344.057 | -0.2428 |
| 778.900 | 778.708 | -0.1915 |
| 964.079 | 963.551 | -0.5277 |
| 1085.869 | 1085.361 | -0.5071 |

Analysis for Energy Calibration:

The channel number in which the Europium peaks are observed is within the same range of values. As it can be observed that the residual energy is extremely small for Cobalt, Cesium and Barium, therefore the calibration function can be considered accurate. In order to be certain of the accuracy I used the Europium bin counts to obtain the potential gamma energy values. Here it can be seen that the residual energies for Europium-152 are extremely small and therefore it establishes the fact that the calibration function is accurate. Data taken in the three days are nearly close to one another establishing the fact that the experiment is pretty consistent. The residual energies for Day 1 and Day 2 show minor variations for the last two gamma values whereas Day 3 has a variation within the range 0.3-0.5.

Data taken on the first day are for a duration of 15-20 mins. Therefore, the precision is within the range of 2-3%. The second day data were taken for a longer time (30mins each) for Europium and Uranium and the precision is in range of 0.5-2.5%. The third day data values were taken for around 35-40 mins and the precision is in the range 0.5-2.0%. It’s observed that as the number of counts increases the precision gets lower. Hence the data taken are more accurate.

Relative Detector Efficiency:

The reference peak:

| N_ref | 19578.5 |
| --- | --- |
| Gamma_ref | 121.00 |
| I_ref | 100 |

The corresponding peaks:

| N_i | I_i |
| --- | --- |
| 2850.0 | 27.9 |
| 6790.0 | 97.9 |
| 1229.0 | 48.0 |
| 1196.9 | 53.5 |

Using the reference peaks and the relative intensities of the rest of the peaks the detector efficiency was calculated.

Day 1:

Table 5: Uranium data values - Day1

| Peak | Channel Number | Gamma energy | Intensity |
| --- | --- | --- | --- |
| Uranium-235 |  |  |  |
| 1446.0 | 873.0 | 143.76 | 10.96 |
| 915.0 | 976.0 | 163.33 | 5.08 |
| 7505.0 | 1093.0 | 185.715 | 57.2 |
| 900.0 | 1195.0 | 194.94 | 0.63 |
| Uranium-238 |  |  |  |
| 432.0 | 5313.0 | 1001.7 | 0.838 |
| 261.0 | 4099.0 | 766.36 | 40 |

Day 2:

Table 6: Uranium data values - Day2

| Peak | Channel Number | Gamma energy | Intensity |
| --- | --- | --- | --- |
| Uranium-235 |  |  |  |
| 3567.0 | 873.0 | 143.76 | 10.96 |
| 2307.0 | 976.0 | 163.33 | 5.08 |
| 19068.0 | 1093.0 | 185.715 | 57.2 |
| 2090.0 | 1195.0 | 194.94 | 0.63 |
| Uranium-238 |  |  |  |
| 1140.0 | 5313.0 | 1001.7 | 0.838 |
| 672.0 | 4100.0 | 766.36 | 40 |

Day 3:

Table 7: Uranium data values - Day3

| Peak | Channel Number | Gamma energy | Intensity |
| --- | --- | --- | --- |
| Uranium-235 |  |  |  |
| 2404.0 | 872.0 | 143.76 | 10.96 |
| 1494.0 | 975.0 | 163.33 | 5.08 |
| 12223.0 | 1091.0 | 185.715 | 57.2 |
| 1434.0 | 1194.0 | 194.94 | 0.63 |
| Uranium-238 |  |  |  |
| 711.0 | 5308.0 | 1001.7 | 0.838 |
| 434.0 | 4095.0 | 766.36 | 40 |

The Uranium sample contained both isotopes U-235 and U-238. In the low energy spectrum the energy peaks for U-235 was observed and in the high energy spectrum the energy peaks for U-238 was observed. For U-238 I chose the peaks with the highest number of counts for lower precision.

Efficiency calculated:

Day 1:

Table 8: Efficiency calculated for Eu and Ur-Day1

|  | Europium | Uranium |
| --- | --- | --- |
| Efficiency | 0.52184 +/- 0.00094 | 0.73593 – U235 |
|  | 0.35427 +/- 0.00545 | 0.69094 – U235 |
|  | 0.13077 +/- 0.00048 | 0.65153 – U235 |
|  | 0.11418 +/- 0.00041 | 0.11135 – U238 |
|  | 0.10445 +/- 0.00104 | 0.12643– U238 |

Day 2:

Table 9: Efficiency calculated for Eu and Ur - Day2

|  | Europium | Uranium |
| --- | --- | --- |
| Efficiency | 0.52094 +/- 0.00086 | 0.72825 – U235 |
|  | 0.35798 +/- 0.00550 | 0.68474 – U235 |
|  | 0.13041 +/- 0.00034 | 0.63693 – U235 |
|  | 0.10768 +/- 0.00023 | 0.10609 – U238 |
|  | 0.10212 +/- 0.00089 | 0.11426 – U238 |

Day 3:

Table 10:Efficiency calculated for Eu and Ur - Day3

|  | Europium | Uranium |
| --- | --- | --- |
| Efficiency | 0.51462 +/- 0.00084 | 0.70790 – U235 |
|  | 0.36131 +/- 0.00555 | 0.66745 – U235 |
|  | 0.12720 +/- 0.00031 | 0.62293 – U235 |
|  | 0.10050 +/- 0.00019 | 0.11179 – U238 |
|  | 0.13840 +/- 0.00117 | 0.04734 – U238 |

Isotopic Abundance:

Day 1:

The function is:

y_le=2.186*10-6x2-2.970*10-3x+1.118

y_he=-5.603*10-8x2+3.498*10-5x+1.325*10-1

Figure 5: Energy calibration – low energies and high energies

Table 11: Isotopic ratio and percentage value of U-235 with respect to U-238

|  | Isotopic ratio | Percentage value U-235 |
| --- | --- | --- |
|  | 0.007146 | 0.0070957 |

Day 2:

The function is:

y_le=2.083*10-6x2-2.863*10-3x+1.097

y_he=-3.869*10-8x2-3.367*10-5x+1.112*10-1

Figure 6: Efficiency calibration function – low energies and high energies

Table 12:Isotopic ratio and percentage value of U-235 with respect to U-238

|  | Isotopic ratio | Percentage value U-235 |
| --- | --- | --- |
|  | 0.00655 | 0.00661 |

Day 3:

The function is

y_le=1.873*10-6x2-2.643*10-3x+1.049

yhe=1.325*10-7x2-3.962*10-5x-6.082*10-2

Figure 7: Energy calibration function – low energy and high energy

Table 13: Isotopic ratio and percentage value of U-235 with respect to U-238

|  | Isotopic ratio | Percentage value U-235 |
| --- | --- | --- |
|  | 0.007252 | 0.007200 |

The Actinium series and the Uranium series are the decay chain of Uranium 235 and Uranium 238 respectively. In the energy spectrum of Uranium it can be seen that the energy peaks correspond to the energy of the emitted particles. Significant peaks can be identified that coincides with Thorium in the Actinium series.

In natural uranium, U-235 is ~0.72% and U-238 is ~99.28%. Therefore, the ratio is around 0.0072. Values were calculated for two U-235 peaks respective to two U-238 peaks. It was observed that some peaks have a higher background noise and utilizing those peaks gave erroneous values. Hence, I had to consider the first two peaks of uranium-235 and the two peaks of uranium-238. The percentage value for Uranium 235 during all three days are 0.0070, 0.0066, and 0.0072. Hence it can be considered a typical natural isotopic abundance. If the percentage of uranium 235 is less than 0.72% then it can be said that the uranium is depleted. Since the percentage values are within the natural uranium range it can be said that the sample used in this experiment is natural uranium.